## Risk Acceptability Matrix
Hazard Severity Score |  |  | Description
Catastrophic |  | 5 | Life-threatening injury or catastrophic property damage
Critical |  | 4 | Permant injury or serious property damage
Serious |  | 3 | Medically reversible injury or significant property damage
Minor |  | 2 | Limited injury or property damage
Negligible |  | 1 | Minor injury or property damage
Probability of Occurrence Score |  |  | Description
Frequent |  | 5 | Happens with almost every use of the device
Probable |  | 4 | Occurs the majority of times but not with every use
Occasional |  | 3 | Occurs with increased frequency
Remote |  | 2 | More than one occurrence per year but still unlikely
Improbable |  | 1 | Less than one occurrence per year; isolated events
BAR | Broadly acceptable region
ALARP | As Low as Reasonably Practical
INT | Intolerable
Risk Priomity Number (RPN) [Severity x Probability]
| Severity/Prob. | 1 Negligible | 2 Minor | 3 Serious | 4 Critical | 5 Catostrophic
| 5 | 5 | 10 | 15 | 20 | 25
| 4 | 4 | 8 | 12 | 16 | 20
| 3 | 3 | 6 | 9 | 12 | 15
| 2 | 2 | 4 | 6 | 8 | 10
| 1 | 1 | 2 | 3 | 4 | 5

## Hazard Analysis
# | System | AST Team Member Assignment | Hazard Description | AST Hazard Source, Mechanism, & Outcome | Risk before Mitigation Measures | [5 empty] | Risk Elimination or Mitigation Measures | AST Concurrence with MM | Risk After Mitigation Measures | [5 empty] | Corrective Actions | Verification Method? (A, T, D, I) | AST Concurrence with Verification Method? | Verified? (Y/N) | AST Comments | AST Safety Inspection Checklist? (Y/N)
[5 empty] | Severity | AST Severity (before MM) | Likelihood | AST Likelihood (before MM) | Risk Hazard Index | AST RHI (before MM) |  |  | Severity | AST Severity (after MM) | Likelihood | AST Likelihood (after MM) | Risk Hazard Index | AST RHI (after MM)
1 | Wearable HIVE Recording Module |  | Intermittent failure to trasmit EMG packets causes unintended prosthetic movements |  | 1 Negligible |  | 4 Probable |  | 4 |  | 1. Real-time status information will indicate corrupted or dropped packets or degraded link quality 2. Stop movement command will be sent to prosthetic to prevent unintented movements 3. Wireless communication automatically resumes when link is clear |  | 1 Negligible |  | 3 Occasional |  | 3 |  | 1. Antennae may need to be re-positioned
3 |  |  | Repeated failure to transmit EMG packets causes unintended prosthetic movements |  | 1 Negligible |  | 2 Remote |  | 2 |  | 1. Real-time status information will indicate corrupted or dropped packets or degraded link quality 2. Stop movement command will be sent to prosthetic to prevent unintented movements 3. Wireless communication session is aborted |  | 1 Negligible |  | 2 Remote |  | 2 |  | 1. If connected, disconnect neural sensing module from patient 2. Examine error logs and status information 3. If the root cause is determined to be a component failure or cannot be identified, discontinue device from study
4 |  |  | Electrical harm to patient |  | 3 Serious |  | 1 Improbable |  | 3 |  | 1. The neural sensing module, which conencts to the electrodes will be leakage current tested to comply with 14708-1 standards for implantable devices. This is a more rigorous standard than 60601-1 for external devices. |  | 3 Serious |  | 1 Improbable |  | 3 |  | 1. If connected, disconnect neural sensing module from the patient 2. Remove device from patient 3. Discontinue from use in study
|  |  | Patient discomfort due to heat |  | 2 Minor |  | 1 Improbable |  | 2 |  | 1. The neural sensing module is a low power consumption device and therefore extremely unlikely to generate enough heat to cause discomfort. Similar devices from our lab have consumed <40mW. Furthermore, the PCB will be separated from direct body contact by the plastic enclosure. |  | 2 Minor |  | 1 Improbable |  | 2 |  | 1. If connected, disconnect neural sensing module from the patient 2. Remove device from patient 3. Discontinue from use in study
5 |  |  | Battery hazards |  | 4 Critical |  | 1 Improbable |  | 4 |  | 1. The neural sensing module receives power from the NNP power module, which has been tested under investigational device expemption at Case Western Reserve Univeristy. 2. The wearable HIVE recording module will only be used in a supervised laboratory environment and handled by trained study personell. In an abundance of caution, the power module will be charged in between uses when the neural sensing module is disconnected from the patient. |  | 4 Critical |  | 1 Improbable |  | 4 |  | 1. If connected, disconnect neural sensing module from patient 2. Discontinue from use in study
6 |  |  | Device malfunction due to electrostatic discharge |  | 2 Minor |  | 3 Occasional |  | 6 |  | 1. Intan amplifier contains internal diodes to protect from mild ESD events 2. The wearable HIVE recording module will only be used in a supervised laboratory environment and handled by trained study personell |  | 2 Minor |  | 2 Remote |  | 4 |  | 1. If connected, disconnect neural sensing module from patient 2. Inspect neural sensing module, examine error logs and status information 3. If the root cause is determined to be a component failure or cannot be identified, discontinue device from study
7 |  |  | Neural sensing module places mechanical strain on electrode leads |  | 2 Minor |  | 3 Occasional |  | 6 |  | 1. Percutaneous connector system provides strain relief 2. The neural sensing module is a compact and lightweight device. It will only be handled and connected to the patient by trained study personell |  | 2 Minor |  | 1 Improbable |  | 2 |  | 1. Reposition neural sensing module to reduce any tension in adapter cable
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
Note on D1: Hazard Description should include: 1. Source of potential harm 2. Mechanism by which the harm may be caused 3. Worst credible outcome assuming no mitigation measures employed. See AC 437.55-1.
Note on L1: Measures taken to reduce the risk to the public. (Reducing either the severity or probability). Typically design changes, safety devices, warning devices, procedures and training. See AC 437.55-1
Note on F2: The severity of the worst credible effect without any mitigation measures. See AC 437.55-1.
Note on H2: The probability of occurance of the hazard or failure mode without any mitigation measures. See AC 437.55-1.
Note on J2: Combination of severity and probability to determine qualitative risk to the public. Red indicates unacceptable risk. If red, elimination or mitigation actions must be taken to reduce the risk. See AC 437.55-1.
Note on N2: The severity of the worst credible effect or consequence assuming all mitigation measures are employed. See AC 437.55-1
Note on P2: The probability of occurrence of the hazard or failure mode, assuming all mitigation measures are employed. See AC 437.55-1
Note on R2: Combination of severity and probability to determine qualitative risk to the public. Red indicates unacceptable risk. If red, further elimination or mitigation actions must be taken to reduce the risk. See AC 437.55-1.